## Slide 1: Desgastes no par articular da prótese da anca
Célio Pina
Escola Superior de Tecnologia de Setúbal

## Slide 2: O que é o desgaste
uma mudança cumulativa e indesejável das dimensões de um componente devido à remoção gradual de partículas discretas de superfícies em contacto e com movimento relativo, causado predominantemente, por acções mecânicas.
Early Charnley low-friction
arthroplasty showing the
stainless steel femoral head and severe
wear of
the polytetrauorethylene acetabular cup
Dowson, 2001

## Slide 3: Medição do desgaste
Alterações topográficas
Volume
Massa

## Slide 4: Mecanismos de desgaste
Adesão:
Abrasão;
Fadiga superficial;
Triboquímico

## Slide 5: Mecânismos de Desgaste

## Slide 6: Desenvolvimento de novos pares articulares
Implantes e explantes
Testes de desgaste
Testes de configuração simples
Simuladores

## Slide 7: Simuladores
Fig. 1. Leeds II hip joint simulator.
Fig. 2. Schematic of a single test cell.
Fig. 3. Load (a) and movement
(b) cycles of the Leeds

## Slide 8: Próteses com acetabulo de PE
Taxas de desgaste típicas: 0.04-0.38 mm/Ano
38.8-77.6mm3/ano
Partículas esféricas de 1 mm de diâmetro
1011 partículas/ano
100 000 partículas/passo
(Dowson, 2001)

## Slide 9: Próteses PE
Partículas de desgaste do PE
Domímio activo 0.2-8 µm

## Slide 10: Rugosidade
Uma cabeça femoral rugosa induz um aumento do desgaste do UHMWPE
Próteses com acetábulo de PE

## Slide 11: Próteses com acetábulo de PE
Desgaste do UHMWPE
CoCr
(Isaac et al,1986)
(Iasty et al., 1994)
(Dahm & Dearnley, 2000)

## Slide 12: Próteses com acetábulo de PE
(Lancaster et al, 1996)
Pin-on-plate
Vários materiais (CoCr, SS, alumína e zircónia)
“ é a apropriada a utilização de um valor de Ra inferior a 0.01 mm”

## Slide 13: Próteses com acetábulo de PE
(Najjar et al, 2005)
3,5 anos
Cabeças de titânio (Ti6Al4V)

## Slide 14: Desgaste sensível ao manter a rugosidade – filme de passivação
Partículas

## Slide 15
Diminuir o desgaste
Elevada dureza das cabeças
Polimento ( baixa rugosidade)
Materais cerâmicos técnicos
Acetábulo
Cabeça

## Slide 16: Fractura Alumina/Zircónia
Alumina
Zicónia

## Slide 17: Degradação da Zircónia

## Slide 18: Conjugação Tenacidade/dureza
Revestimentos ou tratamento de superfícies

## Slide 19: Engenharia de superfícies
Revisão
Revestimento: TiN
Substrato: Ti6Al4V
(Raimondi & Pietabissa, 2000)

## Slide 20: Engenharia de superfícies
(Fisher et al, 2002)

## Slide 21: Utilização de revestimentos de fase S em cabeça de aço inoxidável Ortron
Dureza 900-1400 HV

## Slide 22: Indentação

## Slide 23: “Scract test”
image of the scratch obtained with a load of 5kg 0.10mT(2).

## Slide 24
Scratch lip height (mm) versus load for Group II coated and uncoated
Ortron 90.
h
d
b
(a1+a2)
W
Mean line
h/2
FWHM

## Slide 25: Tribocorrosão
Wear track length | 8mm
Pins | 8 mm diameter alumina ball
Frequency | 1 Hz
Number of cycles | 3200
Load | 200g
Maximum Hertzian contact pressure (po) | 786MPa
Maximum shear stress ( max ) for Ortron 90 | 244 MPa
% of yield (Ortron 90) | 50%
Table 3 Physical parameters for corrosion-wear test.
Dead weight
Counterweight
2 axis pivot horizontally constrained by inductive transducers
Corrosion-wear cell
Water in at 37ºC
Offset crankshaft
Sliding Table
Reference electrode
Ag/AgCl
Working electrode
(test sample)
Auxiliary electrode Pt
Water
Jacket
Counterbody
Alumina
(8mm diameter)
Brass screw
Haber-luggir capillary
Water out to heater bath
Water Jacket
Water in
at 37ºC
Electrolyte
Solution composition
25 v/v % bovine serum diluted in a solution of 0.89 wt% NaCl pH8

## Slide 27
a) Ortron 90
c) 0.18mT

## Slide 28: Desgaste do PE
Fig. 4.7.1. Six-station reciprocating multidirectional
pin-on-plate wear test machine: a) View of one wear test station;
b) General view of the reciprocating multidirectional pin on plate wear test machine.

## Slide 29: Pin-on-plate
Six-station reciprocating multidirectional pin-on-plate wear test machine:
a)
Pin
Plate

## Slide 30
a) Ortron 90
b) s-phase coated Ortron 90

## Slide 32
Ra =0.01 µm
Ra =0.006 µm
a) Ortron 90
b) s-phase coated Ortron 90

## Slide 33: Próteses Metal/cerâmico
?

## Slide 34: Metal/Metal
Acetabulo CoCr
Cabeça CoCr
Taxa de desgaste
típica = 0.003-00.7 mm/ano
0.1 mm3/ano
Partículas 25-35 nm
1012-1014 partículas /106 ciclos
106-108 partículas/passo

## Slide 35: Metal/Metal
Chronological changes of mean serum chromium
levels in patients with unilateral total hip prosthesis. Although
the mean serum chromium level increased between 6 months
and third year after implantation, it was stable after 3 years
and there were insignificant differences in the mean serum
levels from years 3 to 7 postoperatively (all being approximately
1.6 μg/L).
Maezawa et al, 2010

## Slide 36: Metal/Metal
a. | b.
Figure 2. THR MOM great diameter cup loosening a)anteroposterior view; b) lateral view.
(Pina and Tapadinhas, 2013)

## Slide 37: Metal/Metal
Tribocorrosão

## Slide 38: Cerâmico/Cerâmico
Acetábulo Alumina
Cabeça Alumina
Taxa de desgaste típica
0.002-0.020 mm/ano
0.05 mm3/Ano
Partículas
5-90 nm
(Dowson, 2001)

## Slide 39
Fig. 5. Wear rate comparisons (log scale) of C/standard PE, C-CLPE, M/M
and C/C bearings.
(Essner et al, 2005)

## Slide 40: Cerâmico/Cerâmico
Fig. 3. 32 months from surgery. (a) (above) frontal projection;
(b) lateral projection. Note the sharp ball head fracture and
detachment of a single, big fragment. The rest of ball head is
still articulating in the UHMWPE socket.
Fig. 5. View of ball head
retrieved fragments
(a) and of the reconstituted ball
head and of the UHMWPE socket
(b).
(Piconi et al, 1999)

## Slide 41: Cerâmico/cerâmico
(Shishido, et al, 2006)

## Slide 42: Cerâmico/Metal
Fig. 1. Volume change for ceramic-on-metal
versus metal-on-metal prostheses.
M/M
C/M
(Frikins et al, 2001)
M/M
C/M

## Slide 43: Cerâmico/Metal
Pin and plate wear rates
OCP readings

## Slide 44: Cerâmico/Metal
Ishida et al, 2009

## Slide 45: Cerâmico/Metal
(a) After 200,000 cycles
After 1.0 million cycles
After 3.0 million cycles

## Slide 46: Fim
Celio.pina@estsetubal.ips.pt